2.5 Requirements Manual for “Buy” Parts

2.5.1 Categorise “buy” parts into mechanical and electrical

Electrical BUY Parts

- Microcontroller (MCU)

- Load Cell Connector

- Load Cell Interface Port

- Noise and Low-pass Filter

- Amplifier

- Analog to Digital Converter (ADC) Module

- Cell Excitation Regulator

- Touchscreen Display

- Latching Push Button

- Internal Memory Module

- USB-A Port

- Power Regulator

- USB-C Port

- Battery Charging Circuit

- Battery

- Misc

Total Electrical Buy Parts: 16

Mechanical BUY Parts

- Protective Carrying Case

- Foam

- 1-Panel Frame

- Internal Aluminium Frame

- Misc (Screws, Fasteners, O-rings, Mounting brackets)

Total Mechanical Buy Parts: 5

Total BUY parts = 21

2.5.2 Table for Buy parts

Electrical parts

| No. | Part | Qty. | Requirements |  | Justification |
| --- | --- | --- | --- | --- | --- |
| 1 | Microcontroller (MCU) | 1 | - 5 V operation; ≥16 MHz clock; - 32 kB Flash, 2 kB SRAM; SPI/I²C/UART interfaces; - industrial temperature −40…85 °C. |  | With a committed MCU, deterministic timing for ADC sampling is made possible and steady handling of user interface, force computations, and communication procedures is guaranteed—something non-real-time processors would not be dependable with. |
| 2 | Load Cell Connector | 1 | - Shielded multi-core cable; - compatible with Innovatest load-cell wiring; - PVC/PUR jacket; - −25…80 °C temperature rating; - low-resistance copper conductors. |  | By lowering EMI coupling and guaranteeing a physically secure link under field vibrations, a devoted shielded connector helps to preserve high signal integrity. |
| 3 | Load Cell Interface Port | 1 | - Panel-mount 5-contact connector; - IP67 protection; nickel-plated brass/PA housing; - rated current ≥ 2 A; - screw-lock interface. |  | Applying an industrial M12 interface guarantees IP67 sealing and ensures consistent pin mapping for every Innovatest load cell, therefore minimizing wiring mistakes. |
| 4 | Noise and Low-pass Filter | 1 | - RC/passive low-pass filter; - cutoff frequency 5–20 Hz; - 1–5 % tolerance components; - low-noise conditioning for mV-level signals. |  | The use of passive filter allows for greater stability and reliability than active ones. Ensuring only frequencies below 5-20Hz pass eliminates most high frequency noise from the load cell. |
| 5 | Amplifier | 1 | - Differential instrumentation amplifier; - gain ≥1000×; CMRR ≥100 dB; input offset <50 µV; - input noise density <10 nV/√Hz. |  | With little drift, a precision instrumentation amplifier guarantees correct differential signal amplification that is absolutely necessary for sub-millivolt load-cell outputs. |
| 6 | Analog to Digital Converter (ADC) Module | 1 | - 24-bit delta-sigma ADC; - differential inputs; - SPI interface; - effective resolution ≥18 bits; - input noise <1 µV RMS. |  | For low load-cell output changes and stable digital conversion with reduced quantization noise, a 24-bit ADC offers the necessary resolution. |
| 7 | Cell Excitation Regulator | 1 | - Selectable 5 V / 10 V excitation; - ripple <10 mVpp; load current ≥50 mA; - regulation error <1 %. |  | Consistent load-cell sensitivity and linearity are guaranteed by a controlled 5/10 V excitation, which also avoids measurement variance caused by supply drift. |
| 8 | Touchscreen Display | 2 | - 10.1″ TFT; 1024×600 resolution; 4-wire resistive touch; - ≥300 cd/m² brightness; 5 V supply; - SPI/parallel interface. |  | While a vivid TFT panel guarantees visibility in different lighting conditions during field use, a resistive touchscreen allows for precise control even with gloves. |
| 9 | Latching Push Button | 1 | - Panel-mount latching switch; - IP65 front; ≥1 A @ 24 VDC; - electrical endurance ≥30 000 cycles. |  |  |
| 11 | Internal Memory Module | 1 | - ≥8 MB SPI-NOR Flash; - 3.3 V operation; ≥100k P/E cycles; - data retention ≥20 years. |  | Onboard non-volatile memory offers safe data logging independent of continuous power or external devices. |
| 12 | USB-A Port | 1 | - USB-A female connector; - USB 2.0 compliant; - rated ≥500 mA; - insertion life ≥1500 cycles. |  | A USB-A interface provides universal compatibility for exporting calibration data across service environments. |
| 13 | Power Regulator | 1 | - Linear/DC-DC regulator stage; - 5 V and/or 3.3 V outputs; - ripple <20 mV; - regulation <2 %; - continuous output ≥500 mA. |  | Dedicated low-ripple regulators maintain stable 5 V and 3.3 V rails, therefore preventing noise injection into delicate analog circuitry. |
| 14 | USB-C Port | 1 | - USB-C receptacle; - 5 V input; - mechanical retention ≥8 N; - compliant with USB Type-C geometry. |  | A USB-C charging interface ensures compatibility with modern power sources and guarantees safe current handling. |
| 15 | Battery Charging Circuit | 1 | - Li-Ion CC/CV charging circuit; - 5 V USB-C input; - charge current up to 1 A; - integrated OV/OC/OT protection. |  | A dedicated Li-Ion charging IC ensures safe CC/CV charging behavior and prevents battery degradation or hazardous operation. |
| 16 | Battery | 1 | - Li-Ion battery ≥5000 mAh; - discharge capability ≥2 A; - integrated protection (BMS); - IEC-compliant cell pack. |  | A high-capacity Li-Ion battery ensures sufficient field autonomy, enabling extended operation without external power. |
| 17 | Misc | 1 | - Mixed electronic passives: 1 % metal-film resistors; - X7R ceramic capacitors; - inductors with adequate Q-factor; - insulated copper wire rated ≥60 V and −20…80 °C; - 2.54 mm pin headers or board-to-wire connectors. |  | Industry-standard passives and shielded wiring ensure stable analog performance, low noise, robust insulation, and reliable board-to-board interfacing. |

Mechanical Parts

|  |  |  |  |  | Justification |
| --- | --- | --- | --- | --- | --- |
| 1 | Protective Carrying Case | 1 | - Polypropylene copolymer case; - IP67 rating; - pressure equalization valve; - internal dimensions ≥420×280×150 mm; - impact-resistant hinges and latches. |  | A rugged IP67 case protects the entire system from dust, water, shock, and transport damage during field deployment. |
| 2 | Foam | 1 | - EVA/PU foam inserts; density ≥30 kg/m³; - CNC/laser-cut cavities; - compression-set resistant; - thickness matched to internal case profile. |  | Custom-cut EVA/PU foam provides precise support for each load cell and electronic component, minimizing vibration and mechanical stress. |
| 3 | 1-Panel Frame | 1 | - Aluminium/steel front panel; - 2–3 mm thickness; - anodized or powder-coated; - cut-outs for display, button, LED; - mounting pattern aligned with internal frame. |  | A rigid front panel provides secure mounting for the display and controls, preventing misalignment and improving long-term usability. |
| 5 | Misc | N./a. | - DIN-compliant fasteners (M3–M6); - stainless-steel A2/A4 or zinc-plated steel; - washers and nuts to EN standards; NBR/EPDM O-rings; - steel/aluminium brackets (2–4 mm). |  | Standardized DIN fasteners and brackets ensure repeatable assembly quality and maintain mechanical stability under vibration. |
| 6 | Screws |  |  |  |  |
| 7 | O-Rings |  |  |  |  |
| 8 | Resistors |  |  |  |  |
| 9 | Capacitor |  |  |  |  |